Technical Report Analyses

Marta Topor

12/09/2021


# Comparison of high and low frequecy noise in raw data

The first part of analysis assesses the noise in the data before any preprocessing was applied. This information can be used to assess the quality of recorded data. We can especially pay attention to low frequency noise which may lead to drifts in the data and 50Hz noise which is an important indication of noise especially for Mobita where impedence could not be tested.

## **0.1-2Hz**

**Results**

* Pirate Plot, the box indicates the IQR and median
* Median (IQR) for EasyCap
* Median (IQR) for Mobita
* Wilcoxon test result
* Bayes Factor
* Fligner-Killeen test for homogeneity of variance



There is a significant difference in variance between EasyCap and Mobita for for the Fligner-Killeen test.  
Bonferroni correction x8 as there are 2 tests for each of the 2 frequency bands at 2 time points (2 \* 2 \* 2)

**Median for EasyCap:** 0.0345817(0.420387)  
**Median for Mobita:** 1.3085746(2.6674428)

**Wilcoxon:** V = 11, p = 0.1054688, Bonferroni p = 0.84375, r = 0.5318432 (large)

**Fligner-Killeen:** X2 = 8.8975165, p = 0.0028556, Bonferroni p = 0.0228447

**BF**


BF\_raw\_low$ttest

##   
## Paired Samples T-Test   
## -----------------------------------------------------------------------------   
## Statistic ±% df p   
## -----------------------------------------------------------------------------   
## V1 V2 Bayes factor10 1.366589 4.067429e-5   
## Wilcoxon W 11.00000 0.1054688   
## -----------------------------------------------------------------------------

## **49-51Hz**

**Descriptive statistics, Wilcoxon test and plot:**

* Pirate Plot with IQR and median
* Median for EasyCap
* Median for Mobita
* Wilcoxon test result
* Bayes Factor
* Fligner-Killeen test



There is a significant difference in variance between EasyCap and Mobita for for the Fligner-Killeen test.  
Bonferroni correction x8 as there are 2 tests for each of the 2 frequency bands at 2 time points (2 \* 2 \* 2)

**Median for EasyCap:** -4.7607959(0.3277775)  
**Median for Mobita:** -5.63999(5.3021439)

**Wilcoxon:** V = 30, p = 0.8457031, Bonferroni p = 6.765625, r = 0.0805823 (small)


**Fligner-Killeen:** X2 = 11.0140189, p = 9.0425376^{-4}, Bonferroni p = 0.007234

**BF**

BF\_raw\_high$ttest

##   
## Paired Samples T-Test   
## -----------------------------------------------------------------------------   
## Statistic ±% df p   
## -----------------------------------------------------------------------------   
## V1 V2 Bayes factor10 0.3170711 6.808948e-5   
## Wilcoxon W 30.00000 0.8457031   
## -----------------------------------------------------------------------------

**Raw Freq Plot Log-Transformed 0-60Hz**



Raw\_Freq


# Comparison of noise during and after pre-processing

In this part data will be analysed to compare artefacts which were removed during pre-processing. Then noise metrics SNR and RMS will be compared between the two systems. Lastly, 0.1-2Hz and 49-51Hz power will be compared between the systems again.

## **0.1-2Hz**

**Descriptive statistics, Wilcoxon test and plot:**

* Pirate Plot, the box indicates the IQR and median
* Median for EasyCap
* Median for Mobita
* Wilcoxon test result
* Bayes Factor
* Fligner-Killeen test



There are no significant results in the data.  
Bonferroni correction x8 as there are 2 tests for each of the 2 frequency bands at 2 time points (2 \* 2 \* 2)

**Median for EasyCap:** 5.3057559(0.5738168)  
**Median for Mobita:** 5.6848829(1.6298652)

**Wilcoxon:** V = 18, p = 0.375, Bonferroni p = 3, r = 0.3062127 (moderate)

**Fligner-Killeen:** X2 = 6.8638651, p = 0.0087956, Bonferroni p = 0.0703649

**BF**


BF\_cleaned\_low$ttest

##   
## Paired Samples T-Test   
## -----------------------------------------------------------------------------   
## Statistic ±% df p   
## -----------------------------------------------------------------------------   
## V1 V2 Bayes factor10 0.4233767 4.514927e-5   
## Wilcoxon W 18.00000 0.3750000   
## -----------------------------------------------------------------------------

## **49-51Hz**

**Descriptive statistics, Wilcoxon test and plot:**

* Piate Plot with IQR and median
* Median for EasyCap
* Median for Mobita
* Wilcoxon test result
* Bayes Factor
* Fligner-Killeen test



There are no significant results in the data.  
Bonferroni correction x8 as there are 2 tests for each of the 2 frequency bands at 2 time points (2 \* 2 \* 2)

**Median for EasyCap:** -5.0572005(0.7788671)  
**Median for Mobita:** -5.0522749(1.1597188)

**Wilcoxon:** V = 34, p = 0.5566406, Bonferroni p = 4.453125, r = 0.209514 (small)


**Fligner-Killeen:** X2 = 0.443075, p = 0.505642, Bonferroni p = 4.0451359

**BF**

BF\_cleaned\_high$ttest

##   
## Paired Samples T-Test   
## -----------------------------------------------------------------------------   
## Statistic ±% df p   
## -----------------------------------------------------------------------------   
## V1 V2 Bayes factor10 0.3718477 5.249598e-5   
## Wilcoxon W 34.00000 0.5566406   
## -----------------------------------------------------------------------------

**Raw Freq Plot Log-Transformed 0-60Hz**




Range\_Freq

## **Rejected artefactual trials**

There were three types of trials extracted for the analyses in this study. Stimulus locked trials, response locked trials with correct responses and response locked trials with incorrect responses. Proportion of rejected trials was calculated only for the stimulus-locked trials as the total number of trials was the largest in this type whereas for the response-locked trials, the total number was split (correct/incorrect).

* Pirate Plot
* Median (IQR) for Easy Cap
* Median (IQR) for Mobita
* Wilcoxon test results
* Bayes Factor
* Fligner-Killeen test



Both, the Wilcoxon and the Fligner-Killeen test are significant.  
Bonferroni correction x2 as there are 2 tests looking at artefact rejection in the dataset.

**Median for EasyCap:** 0.5841695(0.6395922)  
**Median for Mobita:** 3.7933191(12.6008434)

**Wilcoxon:** V = 1, p = 0.0039063, Bonferroni p = 0.0078125, r = 0.8541723 (large)

**Fligner-Killeen:** X2 = 6.4155382, p = 0.0113126, Bonferroni p = 0.0226252

**BF**


BF\_art\_rej$ttest

##   
## Paired Samples T-Test   
## -------------------------------------------------------------------------------------------------   
## Statistic ±% df p   
## -------------------------------------------------------------------------------------------------   
## E\_Stim\_Rej\_p M\_Stim\_Rej\_p Bayes factor10 1.880490 1.322688e-5   
## Wilcoxon W 1.000000 0.0039062   
## -------------------------------------------------------------------------------------------------

## **Signal to Noise Ratio (SNR)**

The SNR was also extracted from stimulus-locked trials. A subset of electrodes was chosen which excluded electrodes on the edges of the caps which are generally more prone to noise. The final electrode subset included: F3, Fz, F4, FC5, FC1, FC2, FC6, C3, Cz, C4, CP5, CP1, CP2, CP6, P3, Pz, P4.

* Pirate Plot
* Median (IQR) for Easy Cap
* Median (IQR) for Mobita
* Bayes Factor
* Wilcoxon test results
* Fligner-Killeen test



The results are non-significant.  
Bonferroni correction x4 as there are 2 tests and 2 measures of noise (SNR & RMS).

**Median for EasyCap:** 0.0457509(0.049751)  
**Median for Mobita:** 0.037662(0.037662)

**Wilcoxon:** V = 48, p = 0.0371094, Bonferroni p = 0.1484375, r = 0.6607748 (large)

**Fligner-Killeen:** X2 = 1.6691743, p = 0.1963692, Bonferroni p = 0.7854767

**BF**


BF\_snr$ttest

##   
## Paired Samples T-Test   
## ---------------------------------------------------------------------------------------------   
## Statistic ±% df p   
## ---------------------------------------------------------------------------------------------   
## SNR\_mean\_E SNR\_mean\_M Bayes factor10 3.212671 4.197004e-6   
## Wilcoxon W 48.00000 0.0371094   
## ---------------------------------------------------------------------------------------------

## **Root-Mean Square (RMS)**

The RMS values were also extracted from stimulus-locked trials. A subset of electrodes was chosen which excluded electrodes on the edges of the caps which are generally more prone to noise. The final electrode subset included: F3, Fz, F4, FC5, FC1, FC2, FC6, C3, Cz, C4, CP5, CP1, CP2, CP6, P3, Pz, P4.

* Pirate Plot
* Median (IQR) for Easy Cap
* Median (IQR) for Mobita
* Wilcoxon test results
* Flinger-Killeen test



Fligner-Kileen test is significant for RMS between EasyCap and Mobita.  
Bonferroni correction x4 as there are 2 tests and 2 measures of noise (SNR & RS).

**Median for EasyCap:** 10.804762(1.8340205)  
**Median for Mobita:** 20.796485(12.3880117)

**Wilcoxon:** V = 5, p = 0.0195312, Bonferroni p = 0.078125, r = 0.7252407 (large)

**Fligner-Killeen:** X2 = 12.7360151, p = 3.586811^{-4}, Bonferroni p = 0.0014347

**BF**

BF\_rms$ttest

##   
## Paired Samples T-Test   
## -------------------------------------------------------------------------------------------------   
## Statistic ±% df p   
## -------------------------------------------------------------------------------------------------   
## RMS\_E\_subset RMS\_M\_subset Bayes factor10 1.513125 3.308283e-5   
## Wilcoxon W 5.000000 0.0195312   
## -------------------------------------------------------------------------------------------------

# Comparison of frequency activity between the two systems

## **Theta analysis**



**1. Electrode matching**

Select the electrodes that were identified to have strongest activity for theta band and which are overlapping between the two systems. (most positive activity)  
Select from a subset of the following electrodes only:  
F3, Fz, F4, FC5, FC1, FC2, FC6, C3, Cz, C4, CP5, CP1, CP2, CP6, P3, Pz, P4  
The above subset excludes electrodes located on the edges of the cap which are prone to noise and extreme values.

Overlapping electrodes: Fz, F4, FC1, FC2

**2. Results**

* Pirate Plots
* Median (IQR) for EasyCap
* Median (IQR) for Mobita
* Correlation Plot
* Wilcoxon Test
* Bayes Factor
* Fligner-Killeen test



There are no statistically significant results in the data.  
Bonferroni correction x8 as there are 4 frequency bands and 2 tests each.

**Median EasyCap =** -0.7805815(0.5875046)

**Median Mobita =** 0.0992178(1.0312601)

**Wilcoxon:** V = 12, p = 0.1308594, Bonferroni p = 1.046875, r = 0.4996102 (moderate)

**Fligner-Killeen:** X2 = 1.0309402, p = 0.3099379, Bonferroni p = 2.4795035

**BF**

BF\_theta$ttest

##   
## Paired Samples T-Test   
## -----------------------------------------------------------------------------------------------------   
## Statistic ±% df p   
## -----------------------------------------------------------------------------------------------------   
## theta\_av\_E\_log theta\_av\_M\_log Bayes factor10 0.9502437 9.213837e-5   
## Wilcoxon W 12.00000 0.1308594   
## -----------------------------------------------------------------------------------------------------

## **Alpha analysis**



**Electrode matching**

Select the electrodes that were identified to have strongest activity for theta band and which are overlapping between the two systems. (most positive activity)  
Select from a subset of the following elecrodes only:  
F3, Fz, F4, FC5, FC1, FC2, FC6, C3, Cz, C4, CP5, CP1, CP2, CP6, P3, Pz, P4  
The above subset excludes electrodes located on the edges of the cap which are prone to noise and extreme values.

Overlapping electrodes: Fz, Pz, P4

**2. Results**

* Correlation Plot
* Pirate Plots
* Median (IQR) for EasyCap
* Median (IQR) for Mobita
* Wilcoxon Test
* Bayes Factor
* Fligner-Killeen test



There are no statistically significant results in the data.  
Bonferroni correction x8 as there are 4 frequency bands and 2 tests each.

**Median EasyCap =** -1.1508447(0.8791393)

**Median Mobita =** -0.3762333(1.8190568)

**Wilcoxon:** V = 18, p = 0.375, Bonferroni p = 3, r = 0.3062127 (moderate)

**Fligner-Killeen:** X2 = 1.2295939, p = 0.267486, Bonferroni p = 2.1398882

**BF**

BF\_alpha$ttest

##   
## Paired Samples T-Test   
## -----------------------------------------------------------------------------------------------------   
## Statistic ±% df p   
## -----------------------------------------------------------------------------------------------------   
## alpha\_av\_E\_log alpha\_av\_M\_log Bayes factor10 0.7933778 3.861620e-5   
## Wilcoxon W 18.00000 0.3750000   
## -----------------------------------------------------------------------------------------------------

## **Low Beta Analysis**



**1. Electrode matching**

Select the electrodes that were identified to have strongest activity for theta band and which are overlapping between the two systems. (least positive activity)  
Select from a subset the following elecrodes only:  
F3, Fz, F4, FC5, FC1, FC2, FC6, C3, Cz, C4, CP5, CP1, CP2, CP6, P3, Pz, P4  
The above subset excludes electrodes located on the edges of the cap which are prone to noise and extreme values.

Overlapping electrodes: CP1, CP2, Pz

**2. Results**

* Correlation Plot
* Pirate Plots
* Median (IQR) for EasyCap
* Median (IQR) for Mobita
* Wilcoxcon Test
* Bayes Factor
* Fligner-Killeen test



There are no statistically significant results in the data.  
Bonferroni correction x8 as there are 4 frequency bands and 2 tests each.

**Median EasyCap =** -2.9987862(0.8986518)

**Median Mobita =** -2.4034859(1.3188573)

**Wilcoxon:** V = 6, p = 0.0273438, Bonferroni p = 0.21875, r = 0.6930077 (large)

**Fligner-Killeen:** X2 = 1.7232285, p = 0.189278, Bonferroni p = 1.514224

**BF**

BF\_lbeta$ttest

##   
## Paired Samples T-Test   
## -----------------------------------------------------------------------------------------------------   
## Statistic ±% df p   
## -----------------------------------------------------------------------------------------------------   
## lbeta\_av\_E\_log lbeta\_av\_M\_log Bayes factor10 1.914701 1.177940e-5   
## Wilcoxon W 6.000000 0.0273438   
## -----------------------------------------------------------------------------------------------------

## **High Beta Analysis**



**Electrode matching**

Select the electrodes that were identified to have strongest activity for theta band and which are overlapping between the two systems. (least positive activity)  
Select from a subset of the following electrodes only:  
F3, Fz, F4, FC5, FC1, FC2, FC6, C3, Cz, C4, CP5, CP1, CP2, CP6, P3, Pz, P4  
The above subset excludes electrodes located on the edges of the cap which are prone to noise and extreme values.

Overlapping electrodes: CP1, CP2, Pz, P4

**2. Results**

* Correlation Plot
* Pirate Plots
* Median (IQR) for EasyCap
* Median (IQR) for Mobita
* Wilcoxon Test
* Bayes Factor
* Fligner-Killeen test



There are no statistically significant results in the data.  
Bonferroni correction x8 as there are 4 frequency bands and 2 tests each.

**Median EasyCap =** -3.6691439(0.8376614)

**Median Mobita =** -3.1071446(0.6716112)

**Wilcoxon:** V = 3, p = 0.0097656, Bonferroni p = 0.078125, r = 0.7897065 (large)


**Fligner-Killeen:** X2 = 0.0728059, p = 0.7872944, Bonferroni p = 6.2983551

**BF**

BF\_hbeta$ttest

##   
## Paired Samples T-Test   
## -----------------------------------------------------------------------------------------------------   
## Statistic ±% df p   
## -----------------------------------------------------------------------------------------------------   
## hbeta\_av\_E\_log hbeta\_av\_M\_log Bayes factor10 4.690504 3.990473e-6   
## Wilcoxon W 3.000000 0.0097656   
## -----------------------------------------------------------------------------------------------------

## **Frequency Topographies**





Plots

# Comparison of Event-Related Potentials between the two systems

## **P300**

**Mean Amplitude**

* Correlation Plot
* Pirate Plot
* Median (IQR) for Easy Cap
* Median (IQR) for Mobita
* Wilcoxon test results
* Bayes Factor
* Fligner-Killeen test

 The results are non-significant. Bonferroni correction x12 as there are 2 types of ERPs with 6 tests used for each

**Median EasyCap =** 2.397499(1.6327702)

**Median Mobita =** 1.6918165(2.387292)

**Wilcoxon:** V = 31, p = 0.7695312, Bonferroni p = 9.234375, r = 0.1128152 (small)

**Fligner-Killeen:** X2 = 1.531328, p = 0.2159132, Bonferroni p = 2.5909589


**BF**

BF\_P300$ttest

##   
## Paired Samples T-Test   
## ---------------------------------------------------------------------------------   
## Statistic ±% df p   
## ---------------------------------------------------------------------------------   
## Pz\_E Pz\_M Bayes factor10 0.3267930 6.437976e-5   
## Wilcoxon W 31.00000 0.7695312   
## ---------------------------------------------------------------------------------

**Latency**

* Pirate Plot
* Median (IQR) for Easy Cap
* Median (IQR) for Mobita
* Wilcoxon test results
* Fligner-Killeen test



The results are non-significant. Bonferroni correction x12 as there are 2 types of ERPs with 6 tests used for each

**Median EasyCap =** 345.70313(37.5976575)

**Median Mobita =** 342.77344(66.8945275)

**Wilcoxon:** V = 21, p = 0.7262861, Bonferroni p = 8.7154338, r = 0.1622214 (small)

**Fligner-Killeen:** X2 = 1.6883725, p = 0.1938156, Bonferroni p = 2.3257877


**BF**

BF\_P300\_lt$ttest

##   
## Paired Samples T-Test   
## -------------------------------------------------------------------------------------   
## Statistic ±% df p   
## -------------------------------------------------------------------------------------   
## Time\_E Time\_M Bayes factor10 0.3175535 6.789447e-5   
## Wilcoxon W 21.00000 0.7262861   
## -------------------------------------------------------------------------------------

**Amplitude**

* Pirate Plot
* Median (IQR) for Easy Cap
* Median (IQR) for Mobita
* Wilcoxon test results
* Bayes Factor
* Fligner-Killeen test



The results are non-significant. Bonferroni correction x14 as there are 2 types of ERPs with 7 tests used for each

**Median EasyCap =** 2.5396665(1.7545155)

**Median Mobita =** 2.4892555(3.2208377)


**Wilcoxon:** V = 23, p = 0.6953125, Bonferroni p = 8.34375, r = 0.1450481 (small)

**Fligner-Killeen:** X2 = 2.8537179, p = 0.0911629, Bonferroni p = 1.0939547

**BF**

BF\_P300\_amp$ttest



##   
## Paired Samples T-Test   
## -----------------------------------------------------------------------------------   
## Statistic ±% df p   
## -----------------------------------------------------------------------------------   
## Amp\_E Amp\_M Bayes factor10 0.3279011 6.398571e-5   
## Wilcoxon W 23.00000 0.6953125   
## -----------------------------------------------------------------------------------

**P300 Plots**  


## **ERN**

**Mean Amplitude**

* Correlation Plot
* Pirate Plot
* Median (IQR) for Easy Cap
* Median (IQR) for Mobita
* Wilcoxon test results
* Bayes Factor
* Fligner-Killeen test



The results are non-significant. Bonferroni correction x12 as there are 2 types of ERPs with 6 tests used for each

**Median EasyCap =** -2.685649(5.464662)

**Median Mobita =** -0.894613(1.61772)

**Wilcoxon:** V = 1, p = 0.0078125, Bonferroni p = 0.09375, r = 0.8490334 (large)

**Fligner-Killeen:** X2 = 1.0369341, p = 0.3085356, Bonferroni p = 3.7024272


**BF**

BF\_ERN$ttest

##   
## Paired Samples T-Test   
## -----------------------------------------------------------------------------------   
## Statistic ±% df p   
## -----------------------------------------------------------------------------------   
## dif\_E dif\_M Bayes factor10 6.963558 5.922718e-7   
## Wilcoxon W 1.000000 0.0078125   
## -----------------------------------------------------------------------------------

**Peak Latency**

* Pirate Plot
* Median (IQR) for Easy Cap
* Median (IQR) for Mobita
* Wilcoxon test results
* Bayes Factor
* Fligner-Killeen test



The results are non-significant. Bonferroni correction x12 as there are 2 types of ERPs with 6 tests used for each

**Median EasyCap =** 41.01563(15.625)

**Median Mobita =** 5.85938(29.29688)

**Wilcoxon:** V = 36, p = 0.0142662, Bonferroni p = 0.1711942, r = 0.8703065 (large)

**Fligner-Killeen:** X2 = 0.3146367, p = 0.5748488, Bonferroni p = 6.898185


**BF**

BF\_ERN\_lt$ttest

##   
## Paired Samples T-Test   
## -------------------------------------------------------------------------------------   
## Statistic ±% df p   
## -------------------------------------------------------------------------------------   
## Time\_E Time\_M Bayes factor10 17.53392 3.319562e-6   
## Wilcoxon W 36.00000 0.0142662   
## -------------------------------------------------------------------------------------

**Peak Amplitude**

* Pirate Plot
* Median (IQR) for Easy Cap
* Median (IQR) for Mobita
* Wilcoxon test results
* Fligner-Killeen test



The results are non-significant. Bonferroni correction x12 as there are 2 types of ERPs with 6 tests used for each

**Median EasyCap =** -3.95719(7.964275)

**Median Mobita =** -4.557438(0.977197)

**Wilcoxon:** V = 23, p = 1, Bonferroni p = 12, r = 0.019745 (small)


**Fligner-Killeen:** X2 = 1.4818293, p = 0.2234885, Bonferroni p = 2.6818624

**BF**

BF\_ERN\_amp$ttest

##   
## Paired Samples T-Test   
## -----------------------------------------------------------------------------------   
## Statistic ±% df p   
## -----------------------------------------------------------------------------------   
## Amp\_E Amp\_M Bayes factor10 0.3246077 5.010149e-5   
## Wilcoxon W 23.00000 1.0000000   
## -----------------------------------------------------------------------------------


**ERN Plots**



Plots

## **Investigating why the P300 is shifed frontally**

**EasyCap and Mobita P300 activity at Fz and Pz**  
EasyCap is in black and Mobita is in red.




P300 waves

Here the activity looks stronger for Fz in Mobita in comparison to Fz in EasyCap and in comparison to Mobita activity at Pz.

**SNR distribution across different areas of electrodes**  
The electrodes will be divided into 3 areas and the mean SNR values for those areas will be compared using the Wilcoxon and Fligner-Killeen tests.

Frontal: Fp1, F2p, F7, F3, Fz, F4, F8,  
Central: FC5, FC1, FC2, FC6, C3, Cz, C4, CP5, CP1, CP2, CP6  
Posterior: P7, P3, Pz, P4, P8, O1, O2



The results from Wilcoxon test are significant for the posterior electrodes. All other tests are non-significant.  
Because there are 4 other tests looking at noise in the data, p-value are be multiplied by 10.  
Bonferroni correction x10.

**Frontal electrodes**  
EasyCap Median(IQR): 0.0472709(0.0540027)  
Mobita Median(IQR): 0.0264437(0.0202856)  
Wilcoxon: 47, p-value = 0.0488281, Bonferroni p-value: 0.4882812, r = 0.6285419 (large)  
Fligner-Killeen: 0.4554226, p-value = 0.499771, Bonferroni p-value: 4.9977105  
**BF**

BF\_snr\_front$ttest

##   
## Paired Samples T-Test   
## ---------------------------------------------------------------------------------------------------   
## Statistic ±% df p   
## ---------------------------------------------------------------------------------------------------   
## SNR\_frontal\_e SNR\_frontal\_m Bayes factor10 1.996629 8.680177e-6   
## Wilcoxon W 47.00000 0.0488281   
## ---------------------------------------------------------------------------------------------------

**Central electrodes**  
EasyCap Median(IQR): 0.0433928(0.0497122)  
Mobita Median (IQR): 0.0349156(0.0233167)  
Wilcoxon: 46, p-value = 0.0644531, Bonferroni p-value: 0.6445312, r = 0.596309 (large)  
Fligner-Killeen: 0.3278975, p-value = 0.5668998, Bonferroni p-value: 5.6689976  
**BF**

BF\_snr\_cent$ttest

##   
## Paired Samples T-Test   
## ---------------------------------------------------------------------------------------------------   
## Statistic ±% df p   
## ---------------------------------------------------------------------------------------------------   
## SNR\_central\_e SNR\_central\_m Bayes factor10 1.957348 1.010133e-5   
## Wilcoxon W 46.00000 0.0644531   
## ---------------------------------------------------------------------------------------------------

**Posterior electrodes**  
EasyCap Median (IQR): 0.0679594(0.0386814)  
Mobita Median (IQR): 0.0307444(0.0341736)  
Wilcoxon: 54, p-value = 0.0039063, Bonferroni p-value: 0.0390625, r = 0.8541723 (large)  
Fligner-Killeen: 0.091518, p-value = 0.7622562, Bonferroni p-value: 7.6225617  
**BF**

BF\_snr\_post$ttest

##   
## Paired Samples T-Test   
## -------------------------------------------------------------------------------------------------------   
## Statistic ±% df p   
## -------------------------------------------------------------------------------------------------------   
## SNR\_posterior\_e SNR\_posterior\_m Bayes factor10 32.43502 1.608480e-6   
## Wilcoxon W 54.00000 0.0039062   
## -------------------------------------------------------------------------------------------------------

## **Number of trials included in analyses**

**Response-locked epochs**

Mean (SD): 27.5(18.9340962)

**Stimulus-locked epochs**

Mean (SD): 467.8(90.5805964)